## Notes
Notes:
Data filled in General and conclusion Sections will not upload in TIIMES. Please use these sections to keep data ready for copy paste.
Data in Item Description columns of Accepted Quantity and cumulative Accepted quantity must be verified thoroughly before uploading in TIIMES.
Pending activity, Areas of concern & Details of measuring instruments sheets are applicable for IVR.
Select options from drop down list for unit of measurement (Item Description) & certifying body of measuring instruments (Details of measuring instruments).
Do not modify this format.
To start next row in cell itself use below shortcut: press 'alt + Enter' keys together. Alt + Enter

## General
Note: - Data filled in this sheet will not upload in TIIMES
Title | Details
Control Number
SAP number | 8117357332
Notification Number | Email Dtd.28.07.2023
Project Name & Location | SCII_4
Client Name | Endress+Hauser (India) Automation Instrumentation Pvt. Ltd.
End User Name | Endress+Hauser (India) Automation Instrumentation Pvt. Ltd.
DEC / PMC / EPC Name | Not Applicable
DEC / PMC / EPC Number | Not Applicable
Vendor Name | M/s. Cold Weld Engineers Pvt. Ltd., Vasai Road
PO Number (Client PO on Vendor) | A10/1018631456 Dtd.11.05.2023
Sub Vendor | Not Applicable
PO Number (Vendor PO on Sub Vendor) | Not Applicable
Inspection Location | Gala no.- A Cold Weld Estate, 22-23 Gokhivare Village, Sativali Road, Near Forest Range Office, Vasai (East).
Inspection Date(s) | 03.08.2023 & 04.08.2023

## Item Description
Po Item No. | Item Code | Item Description | Unit | Po Quantity | Offered Quantity | Accepted Quantity (Please verify thouroghly before using this Column for IRN) | Cumulative Accepted Quanitity (for IRN Only)
010 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILT3008,E+H Hard Marking - 18X5P | Piece | 1 | 1 | 1 | 1 | number
020 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILZT3001,E+H Hard Marking - 18X5R | Piece | 1 | 1 | 1 | 1 | meter
030 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILZT3002,E+H Hard Marking - 18X5T | Piece | 1 | 1 | 1 | 1 | km
040 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILT3003,E+H Hard Marking - 18X5W | Piece | 1 | 1 | 1 | 1 | meter (number)
050 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILZT3019,E+H Hard Marking - 18X5X | Piece | 1 | 1 | 1 | 1 | km (number)
060 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILT3013,E+H Hard Marking - 18X5Y | Piece | 1 | 1 | 1 | 1 | each
070 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILT3024,E+H Hard Marking - 18X60 | Piece | 1 | 1 | 1 | 1 | Piece
080 | 71405586 | Bypass Chamber, 2"/DN50, FAX42, Tag No. 20ILT3012,E+H Hard Marking - 18X61 | Piece | 1 | 1 | 1 | 1 | Test sample
AU - All Unit
lot
Set
Running Meter
kg
metric ton (tonne)
ton
cubic millimetre
cubic centimeter
cubic meter
cubic inch
cubic foot
mm
cm
in
foot
mile
yard
liter
kl
cl
ml
g
lb
oz
Sq. mm
Sq. cm
Sq. meter
Sq. in
Sq. foot

## Reference Documents
Sr. No. | Document Name | Client Document Number | Vendor Document Number | Approval Status
1 | PR/MR/ARM | Not Applicable | Not Applicable | Not Applicable
2 | QAP/ITP/QCP | QAP_Bypass_01 Dtd.06-02-2023 (Standard QAP For Bypass Chamber) | Not Applicable | Approved & Provided By E+H
3 | Drawing(s) | 2023--VMEP-87-I-XD-00072_01 Rev C1 | Not Applicable | Provided By E+H in Code 1
4 | Procedure(s) | Not Applicable | Not Applicable | Not Applicable
5 | Datasheet(s) | Not Applicable | Not Applicable | Not Applicable
6 | Others (Specify) | Not Applicable | Not Applicable | Not Applicable

## Stages Witnessed
Stages Witnessed
QAP clause no. 3: As per QAP sampling plan & client provided drawing requirements, dimensional Inspection for 100% qty. of bypass chambers per variation in C-C distance was witnessed and dimensions found within tolerance limit.
QAP clause no. 3: As per QAP sampling plan & client provided drawing requirements, dimensional Inspection for offered bypass chamber accessories (Fasteners & Gaskets) were witnessed and dimensions found within acceptable limit.
QAP clause no. 4: As per QAP sampling plan requirements, Liquid penetrant testing on all fillet & butt weld joints of 10% qty. of randomly selected bypass chambers per variation in C-C distance was witnessed and no relevant indications observed at the time of inspection. (Witnessed Chamber Tag No.: - 20ILZT3001, 20ILZT3002, 20ILT3013, 20ILZT3024)
QAP clause no. 4: Asper drawing & CWE standard procedure requirements, Hydrotest for all offered bypass chambers was witnessed at test pressure of 23 kg/cm2(g) (Actual: 25 kg/cm2(g)) for Tag No 20ILT3008, 20ILZT3001, 20ILZT3002, 20ILT3003 and 12.24 kg/cm2(g) (Actual: 13.5 kg/cm2(g)) for Tag No 20ILZT3019, 20ILT3013, 20ILZT3024 , 20ILT3012 & holding time of 30 minutes - No pressure drop & no leakages observed during testing.
QAP clause no. 5: Radiography films along with RT summary sheet and weld map for 100% butt weld joints of all offered bypass chamber was reviewed and no relevant indications observed.
QAP clause no. 5: Witnessed PMI of 10% Bypass Chambers from offered lot for PM & Weld (20ILT3008, 20ILT3003, 20ILT3012)as per specification -found satisfcatory.
QAP clause no. 7: Physical verification & visual inspection for 100% qty. of bypass chambers per variation in C-C distance & its accessories (Fasteners & Gaskets) was witnessed and found with good workmanship and free from surface / weld defects & deformities at the time of inspection. (Witnessed Chamber Tag No.: - 20ILT3008, 20ILZT3001, 20ILZT3002, 20ILT3003, 20ILZT3019, 20ILT3013, 20ILZT3024 , 20ILT3012)
As per QAP clause no. 7 & P.O. requirements, Tag no. & E+H hard marking no. punched on flange OD of all offered bypass chambers was verified and found in order.
All measuring instruments/ equipment were verified for continued suitability for intended use, proper identification, calibration status, traceability to national standards & found satisfactory.

## Document reviewed
Document reviewed
As per QAP clause no. 1 & 4: 3.1 Copy of Inspection certificate for material traceability of offered bypass chamber along with raw materials test certificates of pipe, flanges, plate, coupling, fasteners & gaskets as submitted & declared by vendor was reviewed and found to meet specification & drawing requirements.
As per QAP clause no.1: Incoming inspection record for raw materials used for offered bypass chambers as declared & submitted by vendor was reviewed and found in order.
As per QAP clause no. 2: Applicable WPS, PQR, WPQ & Filler Wire for offered bypass chambers as declared & submitted by vendor was reviewed and found in order.
As per QAP clause no. 5: PMI test report for offered bypass chambers as declared & submitted by vendor was reviewed & found in order.
Test certificates for DP test material used during testing was reviewed / verified and found in order.
As per QAP clause no. 6: Reviewed Painting report of Bypass Chamber Tag No 20ILT3008, 20ILZT3001 , 20ILZT3019, 20ILT3013 as per approved drawing & Painting schedule No PRD/WI-VME-0 R00 - found in order.

## Details of measuring instrument
Sr. No. | Detail of Equipment(s) / Instrument(s) | Range | ID | Validity till Date | Calibration Certificate Number | NABL / OTHER CERT. BODIES / NABL Traceability (type Yes wherever applicable) | NON-NABL (type Yes wherever applicable)
1 | Vernier Caliper | 0-300 MM | 19312458 | 2024-01-15 | 2023/56648/006 Dtd.16/01/2023 | Yes |  | Yes
2 | Measuring Tape | 0-5 M | 56648/005 | 2024-01-15 | 2023/56648/005 Dtd.16/01/2023 | Yes |  | No
3 | Pressure Gauge | 0-21 kg/cm. sq | 002212486 | 16.01.2024 | 578 Dtd.17/01/2023 | Yes
4 | Pressure Gauge | 0-21 kg/cm. sq | 002212487 | 16.01.2024 | 578 Dtd.17/01/2023 | Yes
5 | Pressure Gauge | 0-42 kg/cm. sq | N025.59-00080 | 15.01.2024 | 2023/56648/012 Dtd. 16.01.2023 | Yes
6 | Pressure Gauge | 0-42 kg/cm. sq | 01604503 | 15.01.2024 | 2023/56648/008 Dtd. 16.01.2023 | Yes
7 | PMI Machine | 808152 | --- | 05.09.2023 | JCC1173/22-23 Dtd. 06.09.2022 |  | Yes

## Areas of concerns
Areas Of Concern
NIL

## Pending Activities
Pending Activities from this visit
NIL

## Conclusion
Note: Data filled in this sheet will not upload in TIIMES
Non Conformities raised / Discrepancies observed during Inspection (If Any, Mention TUV India NCR Number ):
None
Identification of Inspected Item(s):
Inspected bypass chambers identified with single hard stamped (low stress) as " " on chamber flange OD near tag no. punched.
Conclusion (Acceptable / Not Acceptable / Acceptance subject to):
Acceptable.
IRN can be issued for PO Item numbers:
Expenses:
Remarks: